13 June 2007

DOSE DEVIATION DURRING STEREOTACTIC RADIOSURGERY TREATMENTS 

DATE   :  6 April 2006 to 17 April 2007

FACILITY INVOLVED : Beam therapy center

FACILITY PLACE  :  Toulouse University Rangueil hospital

An information note and a report was issued.

DESCRIPTION :

The ASN has been informed on April 20th, 2007 of a deviation between the delivered dose and the prescribed dose to 145 patients treated by stereotaxic radiosurgery from April 6th, 2006 to April 17th, 2007.

This event notification followed the discovery by the manufacturer of a deviation concerning the beam calibration. As a result, the unique dose delivered during a one-session was more important to the prescribed dose, in variable proportions depending to the localisation of the organs and the characteristics of the beams that were used. 

ASN confirmed the dose deviations during a reactive inspection performed on May 3rd, 2007 which showed that an inadequate beam calibration tool has been used.

This event is without known consequences today. An epidemiological survey will be organised to follow up the patients.  

See technical report in the English section of ASN Website www.asn.fr. An additional preliminary report concerning human an organisational issues will be published by the end of July 2007.

CONTACT PERSON FOR FURTHER INFORMATION:

Name:  David Krembel

Nuclear Safety Authority

Address : 6, place du Colonel Bourgoin F-75572 Paris Cedex 12,

Tel +33 1.40.19.88.44

Email: david.krembel@asn.fr
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